## Slide 1: Real-time least-squares pulse fitting for digital spectroscopy
Christopher Crawford, University of Kentucky
SRO VIII workshop, 2021-04-29

## Slide 2: Executive Summary
We present the SLS-RPP algorithm, which can do fully covariant nonlinear least-squares template pulse fitting in real-time on streaming data
For SRO: this enables the application of sophisticated Level-1 trigger logic, for example: low-threshold triggers and energy-summing cluster triggers
General and extensible building blocks for wide array of applications, and Generic runtime-configurable firmware
https://leoburnett.com
arXiv:2012.05937 [physics.ins-det]
Notes: For those of you who are double- or triple- zooming, I learned to do executive summaries at Nuclear Science Advocacy Day today, so here you go:

## Slide 3: Outline
0. Physics background
  Nab: neutron decay correlations
Standard tools of digital spectroscopy (DPP)
  Composition of digital logic building blocks
  Trapezoid filter in two pieces
Sliding Least Squares (SLS) fitter
  as a Finite Impulse Response (FIR) filter
Recursive Piecewise Polynomial (RPP) filter
  implementation in FPGA
https://leoburnett.com
Notes: This talk will be different from others in this workshop: instead of summarizing a broad range of progress, I will focus on explaining one particular algorithm for processing streaming data.
I fear that it is a little elementary for this audience, but it emphasizes the power you can achieve with very simple building blocks.

## Slide 4: Neutron decay
Ee dependence of a and A can separately confirm CVC and the absence of SCC.
[Gardner, Zhang, PRL 86, 5666 (2001), Gardner, hep-ph/0312124]
UCNA Brown (2018)
UCNA (2018)
Notes: This development was done in context of the Nab experiment, which will study neutron decay, not the average decay rate, but it’s dependence on the correlation of particle momenta, energy and spin, of the incoming and outgoing particles. By parity violation the correlations are all proportional to the axial coupling GA, the average rate has been factored out of each expression.

## Slide 5: The Nab experiment
Notes: The electron-neutrino angle can be extracted from the outgoing electron and proton energy.
However, the proton energy is too small to detect, so we measure it’s time-of-flight in a field-expansion spectrometer.

## Slide 6
Notes: Here is a clip of the installation of the 7 m SC solenoid to longitudinalize the proton’s momentum to get it’s total energy from the time-of-flight.

## Slide 7
Event Structure 100-750 keV electron, 100 ps systematic, 2 keV resolution 30 keV (HV) proton, 12-40 ms TOF delay, 10 ns resolution
Data rate 250 Hz coincidence adjacent pixels; accidentals 2.5 kHz singles 256 channel digitizer, 250 MS/s rate, 14 bit ADC = 28 MB/s
Notes: The electrons and protons are detected in an ion-implanted silicon detector with127 segments to filter out real coincidences, which occur along the same magnetic field line. One detector is at -30kV to accelerate the proton for detection.

## Slide 8: NI-5171R Digitizer
128 channels per detector / chassis
Single controller
for both detectors: global trigger and readout
Notes: At the time we developed our DAQ system, NI was one of the only commercial boards with an open FPGA, programmable in LabVIEW, which I like a lot better on FPGA than CPU. The library includes an ADC block on one end a DRAM controller, and a PCI DMA-FIFO on the other, and you can put anything in between. It is nice to see Open-Source FPGA available in the new CAEN digitizers now.

## Slide 9: NI-5171R Digitizer
Single controller
for both detectors: global trigger and readout
Red Pitaya digitizer
Zynq 7010 SoC
4 GB RAM
2 ADC + 2 DAC @ 14 bit, 125 MS/s
open-source software AND firmware!
Readout and trapezoid firmware developed on NI-5171R
(Aaron Jezghani and David Mathews, UK)
Trapezoid filter firmware developed on STEMlab 125-14
(Yuke Wang, University of Kentucky)
Notes: We have also worked with the much less expensive hobbyist board Red Pitaya, which has open-source firmware, and my students have developed readout and trapezoid filters on both boards.

## Slide 10: Event Buffered Digitizers
Notes: We have taken a different approach to readout: instead of triggering individual event buffers, and then validating events for readout,

## Slide 11: Continuous Buffered SRO
Notes: The waveform is continuously buffered in a giant ring buffer, while simultaneously going through the energy and timing filters and discrimination. We avoid level 2 trigger logic by sending a lightweight stream of timing and energy events to global logic in CPU, which can then decide which segments to readout, including waveforms that did not trigger at all. In that sense it is streaming readout, and our approach is to filter down the data as soon as possible to minimize unnecessary network traffic. I will show next that it is possible to put completely generic filters on the FPGA, so that all custom triggering can developed and debugged fully on CPU, like Sergey talked about for CLASS12 yesterday.

## Slide 12: Firmware HDL (Verilog/VHDL)
module hold (clk, in, out); //***** Delay buffer *****
parameter bit = 64; //Size of input and output
parameter delay = 1; //Delay amount in clock cycles
input clk;
input signed [bit-1:0] in;
output signed [bit-1:0] out;
reg signed [bit-1:0] array [0:delay + 1]; //Array to hold inputs for delay
integer i; //Register for the for-loop
always @ (posedge clk) begin //Shifts all elements down in array
once
array [0] = in;
for (i = 0; i <= delay; i = i + 1)
array [i+1] <= array [i];
end
assign out = array[delay + 1]; //Assigns final element to output
endmodule
LabVIEW FPGA Module
Trapezoid Filter Schematic

## Slide 13: Digital FIR filters
E1
E2
t1
t2
E1
E2
t1
t2
Notes: Before describing the Sliding Least Square-Recursive Piecewise Polynomial filter, I would like to describe the standard trapezoid filter in terms of basic building blocks, because may of the pieces are the same. Essentially, the trapezoid averages out charge on peak of a tail pulse minus the exponential decaying baseline to get precise energy resolution even in the case of ballistic deficit, where not all the charge is collected at once.

## Slide 14: Building blocks
basic arithmetic adder, multiplier -> Multiply-Accumulate in DSP48 slice
unit impulse basis function delay (identity) shift register/buffer
unit step const function integration accumulator
properties of convolution
  commutative (function <-> filter)
  associative (build up complex filters in stages)
Notes: For building blocks, we have basic arithmetic: adding can be done in the fabric, but Xilinx FPGA have special DSP slices to multiply efficiently.
Then there is just the delta function, which acts as a basis function on it’s own, or as a delay, when convolved with any input function. The value it picks of back here is output at this time, delayed by this much. In digital logic, this can be implements as a chain of shift registers, or a buffer.
The Heaviside step function represents both instant accumulation of charge, and as an integrator filter, when convolved with another function. It is implemented as an accumulator with a shift register to keep the running total.
The properties of convolution allow the interchange between function and filter (that’s why we flip one function when convolving), and allow us to build up a complex filter one convolution at a time.

## Slide 15: Boxcar filter
Baseline subtract
Notes: Ignoring the bleed resistor, the input voltage will be a step function as charge is collected.
You can average this with a boxcar filter, with a linear output as it crosses over the rising edge and averages more of the pulse. This is the delayed difference of two accumulators: each step you add one sample on the front and subtract it off the back.
We remove the baseline, by subtracting a second boxcar, with some separation to allow some time for charge collection. This gives rise to a trapezoid with a flat top of width the settling time, and rise and fall times for the duration of the average.

## Slide 16: Tail pulse
pole-zero
Notes: Adding back in the bleed resistor, the tail pulse can be straightened out by adding an integrated copy, so that the filter includes delta and an integrating step.

## Slide 17: Trapezoid filter
Notes: Combining these operations, the total filter is both the pole-zero and the double-box car. Convolving these together results in the trapezoid filter, but it is implements as a series of filters in digital logic.

## Slide 18: Trapezoid filter

## Slide 19: Full Trapezoid filter
David Mathews, University of Kentucky, APS DNP talk, April 2019

## Slide 20: Double trapezoid filter
Flat top for precise energy measurement
Low-threshold trigger: long shaping-time to average out noise
Zero-crossing for precise trigger timing
Integrated baseline subtraction (zero net area filter)
threshold
energy
timing

## Slide 21: Motivation · Experiment · Detector · Timing · DAQ · DSP · Results
Digital Filters for Signal Processing
Energy Extraction
Arrival Time Determination
Noise Reduction
Aaron Jezghani, University of Kentucky, PhD Thesis
2020-01-10
UNAM-IF Lunchtime Seminar

## Slide 22: Generalized linear LSQ fitting
Original data
Pseudoinverse of design matrix
Fit Functions: Design Matrix A
Fit parameters
David Mathews, University of Kentucky, APS DNP talk, April 2019
Notes: Now we apply these same principles to least squares pulse fitting.
In generalized linear fits, the design matrix is formed from a bunch of template waveforms and noise you want to reject. We fit the waveform to a linear combination of these templates. This is accomplished by multiplying by the pseudoinverse of the design matrix, which gives you each of the coefficients.

## Slide 23: Add nonlinear parameter t0
Pseudoinverse Fit Functions
Data:
Fit Functions:
Fit using Parameters
Calculate Fit Parameters:
Find Best Fit with
David Mathews, University of Kentucky, APS DNP talk, April 2019
Notes: Now we add the nonlinear fit parameter T0 by essentially performing the linear fit at each time. The matrix or dot products turn into convolutions or sliding dot products as we scan across the waveform. Here you can see all 4 coefficients as a function of T0, and a peak in the template amplitude with all noise amplitudes falling away as the template matches the waveform at just the right spot.
Here you can see a bad fit in orange just before it reaches T0, and the bad fit in red just afterward.
By expanding out the normal expression of chi2, we get an expression suitable for streaming implementation: convolution of waveform pointwise squared with a box filter, and pointwise sandwiching of the small vector of fit parameters around a constant matrix. We minimize this to complete the chi2 fit.

## Slide 24: Generic Filter Design
(
()
(
…and extending beyond:
From Jordanov and Knoll:
Aaron Jezghani, University of Kentucky, PhD Thesis
Notes: The remaining task is to fit a convolution consisting of up to thousands of points on an FPGA.
We do this generically with a piecewise polynomial response function.
We already say how to build up each order of polynomial:
A step function (integrator) filters a step (constant) function into a linear function.
Likewise, it integrates a linear into a quadratic, and so on.
Actually, the numbers you get are the same as Pascale’s triangle for the exact same reason.
We subtract off the integral at the end to restrict the polynomial the interval 1..L
Here is the implementation in digital logic, using accumulators, delays, and multiplies.
Now we only need to transform our piecewise polynomial response function into the basis of the Pochhamer polynomials, and we can perform any generic FIR filter, including those needed for the SLS filter.

## Slide 25: Motivation · Experiment · Detector · Timing · DAQ · DSP · Results
Sliding Least Squares Fitting
Pulse
Sine
Cosine
Constant
Design
Matrix
Fit
Parameters
Weighted
Data
a = A-1β
a = A+β
Aaron Jezghani, University of Kentucky, PhD Thesis

## Slide 26: Motivation · Experiment · Detector · Timing · DAQ · DSP · Results
Comparison with Synthetic Signals
Aaron Jezghani, University of Kentucky, PhD Thesis

## Slide 27: Motivation · Experiment · Detector · Timing · DAQ · DSP · Results
Improved Peak Resolution
Pixel 64E
Pixel 64W
Trapezoid
SLS Fit (FFT)
Trapezoid
SLS Fit (FFT)
keV
keV
keV
keV
Aaron Jezghani, University of Kentucky, PhD Thesis
2020-01-10
UNAM-IF Lunchtime Seminar

## Slide 28: Motivation · Experiment · Detector · Timing · DAQ · DSP · Results
Neutrons | HV On | AFP Off
Neutrons present to decay, so protons are present
Accelerating potential pushes protons to detectable energy (30 keV → 18 keV after dead layer)
Neutron spin points WEST → neutrinos preferentially go EAST → more same-side electron-proton coincidences in WEST detector
Neutrons | HV On | AFP Off
Neutrons present to decay, so protons are present
Accelerating potential pushes protons to detectable energy (30 keV → 18 keV after dead layer)
Neutron spin points EAST → neutrinos preferentially go WEST → less same-side electron-proton coincidences in WEST detector
AFP transport efficiency further reduces neutron density
Neutrons | HV Off | AFP On/Off
Neutrons present to decay, so protons are present
No accelerating potential to provide protons with sufficient energy for detection
Sanity check that any e-p “signal” is not just low-energy electrons from beta decay
Background | HV On | AFP On/Off
No neutrons to decay, so protons are absent
Accelerating potential to provide protons with sufficient energy for detection
Sanity check that any e-p “signal” is not just spurious noise due to HV or other particles being accelerated
Proton Detection in UCNB
Neutrons | HV On | AFP Off
Neutrons | HV Off | AFP On/Off
Neutrons | HV On | AFP On
Bkgd | HV On | AFP On/Off
Microphonics
e-p coincidences!
Probably
Microphonics?
Backscatters
Aaron Jezghani, University of Kentucky, PhD Thesis

## Slide 29: Covariant LSQ Fitter
measure noise spectrum
acquire average template waveforms representing nonlinear variations of signal (pulse-shape)
LSQ inverse of design matrix w/r cross-correlation matrix to project out backgrounds
FIR filter for each response function approximated by piecewise polynomials
calculate X2 in real-time, repeat linear LSQ fit for t0 at minimum X2

## Slide 30: Generic filter design!
Aaron Jezghani, University of Kentucky
Notes: In conclusion, the addition of powerful digital filters into the streaming readout system can be used to form versatile trigger logic, or dramatic reduction in data size without loss of essential information and can all be done in real time at full digitization speed. Here, my student Aaron was playing around with a new filter to convert tail pulses into a cool logo basketball-winning logo (at least usually).